## Slide 1
Introduction

## Slide 2: Who am I
Giacomo Cappon
  Assistant Professor, Department of Information Engineering, University of Padova
  PhD in Information Engineering, curriculum Bioengineering
  Main research interests: signal processing; decision support systems; wearable sensors; digital health; ai for medicine
Contacts
  email: giacomo.cappon@unipd.it
  personal website: https://gcappon.github.io
  public repos: https://github.com/gcappon

## Slide 3: What is the course about?
E-health or digital health: The use of modern Information and Communication Technologies (ICT) to efficiently deliver all kinds of healthcare services
  Examples: EHR, telemedicine systems
Wearables: electronic devices worn close to and/or on the surface of the skin, which detect, analyze and transmit biomedical data, such as heart rate, activity, blood pressure, temperature, etc.
Internet-of-thing (IoT) wearables: Wearables that connect and exchange data with other devices (wearable or not) and systems over the Internet or other communication networks
E-health
Wearables
IoT wearables
Biomedical wearable technologies…

## Slide 4: What is the course about?
Our society is facing some important global health challenges
  Lifestyle-related illnesses driven by sedentary life, consumption of processed and fast foods, alcohol drinking and smoking
    Need to change the paradigm of healthcare, from a mostly reactive system to a proactive system
  Population ageing rises the demand of health services and the healthcare expenditure.
    Need to decentralizing the healthcare
    Covid-19 pandemic called for urgent delivery of healthcare at home
… for healthcare and wellbeing
Biomedical wearable technologies are fundamental to make healthcare scalable and accessible outside the hospital setting and deliver personalized and preventive health services

## Slide 5: Theory program
The theoretical part of the course will cover the following:
  Wearable sensor technologies
  Health design thinking
  Elements of UI/UX Design
  Evaluation of app usability
  Fundamentals of network architectures and the HTTP communication protocol
  The representational state transfer (REST) paradigm
  Application programming interfaces (APIs)
  Elements of network security
  Authentication protocols
  Regulation for health data management

## Slide 6: Lab program
The laboratorial part of the course will cover the following:
  The Git version control system
  Dart programming language and the Flutter framework
  Building mobile/web applications in Flutter
  User Interface design
  Navigation
  State management
  Basic data persistence
  Networking
  Integrating AI models in mobile applications
Workshop…

## Slide 7: Lab program - The workshop
2-hours workshop where we will set up a fake clinical trial aimed at evaluating the cardiovascular tone in a cohort of healthy subjects using an experimental algorithm
The fastest 5 students to submit the right answer will get a +1 to the final evaluation.
1. Trial setup
2. Data collection using a wearable and an app
3. Data preparation
4. Data analysis
5. Trial outcome
Together
Individually

## Slide 8: Prerequisites & End of class goals
Prerequisites
  Basic:
    Propension for programming, Python
  Preferable:
    Object-oriented programming
End of class goals
  Knowing how wearable devices communicates to the end-user
  Designing digital solutions to enhance healthcare and well-being
  Developing a fully functioning mobile application

## Slide 9: The exam
Written test: Theory exam
  4 open questions
  Duration: 90 min
  Max vote: 20/30
Project: Development of an application that integrates and elaborates data from a smartwatch. To be done in groups of 3 students.
  Duration of the discussion: 45 min (25 min presentation + 20 min Q&A)
  Max vote: 12/30
Final mark: sum of the votes of the written test and the project.
Dates of the written test
First attempt: 18 Jun 2026, 14:00 (Room Me+Ne)
Second round: 07 Jul 2026, 14:00 (Room De)
Third round: 28 Aug 2026, 14:00 (Room De)
Fourth round: 01 Feb 2027, 14:00 (Room Ee)

## Slide 10: Project: What’s the spirit here
The goal is learning “How to do SOMETHING ORIGINAL using wearable devices”
What I DO expect from you
  To use your imagination to create that SOMETHING
  To learn to build something from scratch
  To use code built by others
  To work together
  To use Google and StackOverflow
What I DO NOT expect from you
  To create the new Google
  To reinvent the wheel (copying others code is fine)
  To create extra complicated UI with animations and stuff

## Slide 11: Project structure
The project consists of building an app for iOS or Android that collects user data from a wearable device through Web APIs, visualizes them, and does some tricks with them.
Core functionalities:
  Correct app flow
  User authentication and management
  Data collection
  Data visualization and presentation
Additional fuctionalities  It’s up to you! Some examples:
  Run some analysis on data and provide suggestions to the user
  Implement some literature algorithm
  …

## Slide 12: Grading
What are the grading criteria
  Compliance to core functionalities
  Originality of the additional feature
  Quality of teamwork
How much these criteria matter?
  Core functionalities: 40%
  Additional app functionalities: 50%
  Quality of teamwork: 10%
Chart: 
| Core functionalities | Additional functionalities | Quality of teamwork
Sales | 40 | 50 | 10

## Slide 13: Choose your target
Your app must represent a solution for the current SDGs (Sustainable Development Goals) targets
SDG are 17 topics representing the current big challenges humanity must face and solve

## Slide 14: Choose your target
These are the target you can choose from (see the elearning page for more details):
  SDG 2: Target 2.2 (By 2030, end all forms of malnutrition, including achieving, by 2025, the internationally agreed targets on stunting and wasting in children under 5 years of age, and address the nutritional needs of adolescent girls, pregnant and lactating women and older persons)
  SDG 2: Target 2.3 (By 2030, double the agricultural productivity and incomes of small-scale food producers, in particular women, indigenous peoples, family farmers, pastoralists and fishers, including through secure and equal access to land, other productive resources and inputs, knowledge, financial services, markets and opportunities for value addition and non-farm employment)
  SDG 3: Target 3.4 (By 2030, reduce by one third premature mortality from non-communicable diseases through prevention and treatment and promote mental health and well-being)
  SDG 3: Target 3.5 (Strengthen the prevention and treatment of substance abuse, including narcotic drug abuse and harmful use of alcohol)
  SDG 4: Target 4.7 (By 2030, ensure that all learners acquire the knowledge and skills needed to promote sustainable development, including, among others, through education for sustainable development and sustainable lifestyles, human rights, gender equality, promotion of a culture of peace and non-violence, global citizenship and appreciation of cultural diversity and of culture’s contribution to sustainable development)
  SDG 8: Target 8.9 (By 2030, devise and implement policies to promote sustainable tourism that creates jobs and promotes local culture and products)

## Slide 15: Data source
Data that you will use are collected by me using a Fitbit Versa 2
Data are available starting from 09/02 and will be collected up to 30/11
The following data “types” are included(see mor details in the Appendix):
  Calories
  Distance
  Exercise sessions
  Heart rate
  Resting heart rate (Estimated)
  Sleep
  Steps
Your project must use at least 1 type of data

## Slide 16: How to get data - The IMPACT backend
https://impact.dei.unipd.it/bwthw/
Your app
HTTP requests/responses
You will be provided with credentials (different for each group) and learn how to get data during the lab lessons

## Slide 17: Optional tutoring activities
Need help? They are here to cover your back with answers and optional tutoring activities
Calendar of the optional tutoring activities
Luca Cossu
luca.cossu@phd.unipd.it
Alberto Gastaldello
alberto.gastaldello.2@phd.unipd.it
OR

## Slide 18: Discord server
Join the Discord server to get in touch and ask questions
Link: https://discord.gg/j76A3D27bm

## Slide 19
Wearables

## Slide 20: Wearable devices

## Slide 21: Wearable devices
A wearable device is a device that can be used for sensing physiological parameters or signals which generally include two components:
  A sensor that measures the physical quantity of interest (measurand)
  An embedded processing unit that does some initial processing of the measurements and either stores them in a memory or sends them to another device
Sensor
Embedded processing unit
Receiver device / Gateway device (for IoT devices)

## Slide 22: Definitions of sensor
“A sensor generally refers to a device that converts a physical measure into a signal that is read by an observer or by an instrument” (Chen et al., 2012)
Measurand

## Slide 23: Classification of sensors based on their interaction with the measurand
Contact sensor: sensors requiring physical contact with the measurand.
Their deployment obviously perturbs the state of the subject to some degree.
  Comfort and biocompatibility are important considerations
  Response time determined by the speed at which the measurand is transported to the measurement site.
  Poor contact can result in signal noise and the introduction of signal artifacts.
  Invasive: sensors inserted into human body
  Minimally-invasive: patch-type devices on the skin that monitor interstitial fluids
  Non-invasive: sensors that simply have contact with the skin without effect

## Slide 24: Classification of sensors based on their interaction with the measurand
Non-contact sensors: do not require direct contact with the measurand.
Minimum perturbation of the subject
Generally used in ambient sensing applications
Example: infrared thermometer
Sample based sensors: require an invasive collection of a sample of the person to be monitored by a human or automated sampling system.
The collected samples are analyzed using sensors or laboratory-based analytical instrumentation
Common approach in healthcare (e.g. monitoring cholesterol)

## Slide 25: Classification of sensors based on their sensing technique
Mechanical sensors
Electrical sensors
Optical sensors
Chemical sensors

## Slide 26: Mechanical sensors
Mechanical sensors detect some form of mechanical deformation of a material, in response to the measurand (input), and translate it to an electrical, optical, magnetic or thermal signal (output).
Electromechanical sensors:
Types of electromechanical sensing techniques:
  Piezoresistive sensors
  Capacitive sensors
  Iontronic sensors
  Piezoelectric sensors
  Measurand
  Mechanical deformation
  Electrical signal

## Slide 27: Electrical sensors
Electrical sensors: on-skin electrodes that measure a change in electrical resistance or the electrical charge of the skin.
Good electrical contact with skin is required.
Wet electrodes: solid conductive pad interfaced to the skin via an electrolyte gel that minimizes the impedance of skin by hydrating it and forming a conformal electrical contact with its textured surface.
Dry electrodes: direct contact with the skin, sensible to movement artifacts.
Skin electrodes allow to collect high-fidelity measurements of a broad range of physiologically relevant biopotentials:
  electrocardiogram (ECG)
  electroencephalogram (EEG)
  electromyogram (EMG)
  electrooculogram (EOG)
  electroretinogram (ERG)
  galvanic skin response (GSR), also known as skin impedance or electrodermal activity (EDA)

## Slide 28: Optical sensors
Electro-optical sensors detect light and convert it to an electrical signal
  Photoconductive devices convert a change of incident light into a change of resistance.
  Photovoltaics (solar cells) convert an amount of incident light into an output voltage.
  Photodiodes and phototransistor convert an amount of incident light into an output current.
Operating principle of wearable optical sensors: a light source introduces light into the body and based on how much the light is absorbed or reflected, measurements related to the body tissue are derived.
The wavelength of light sources can range from UV to deep infrared, depending on needed penetration depth and the absorption peak for the relevant sensing application.
Commonly used for monitoring of heart rate and blood oxygenation.

## Slide 29: Chemical sensors
They quantify chemical markers, or analytes, in biological fluids and convert this information in an optical (optochemical sensor) or electrical signal (electrochemical sensor).
Require direct chemical interaction with the biomarker (contact sensors).
Wearable chemical sensors use non-invasive or minimally-invasive techniques to measure analytes in alternative biofluids instead of invasive blood sampling.
  Biofluids secreted by the body (saliva, sweat, tears)
  Interstitial fluid reached through minimally-invasive sensors or by reverse iontophoresis

## Slide 30: Sensor characteristics
To truly understand sensors, and how sensors that measure the same measurand can differ, it is necessary to understand sensor performance characteristics, which are reported in specifications (specs) or datasheet.
  Systematic characteristics: those which can be exactly quantified by mathematical or graphical means;
  Statistical characteristics: those which cannot be exactly quantified;
  Dynamic characteristics: those characterizing the ways in which an element responds to sudden input changes.

## Slide 31: Transfer function
A function f that describes the relationship between the measurand and the output signal (electrical or optical).
Represented in specs as a table of values, a graph, or a mathematical formula.
Typically the actual transfer function is complex  the transfer function is approximated by a simpler linear or non-linear function, F(.), which can be mathematically expressed as a formula.
e = approximation error
Output electrical (or optical) signal
Measurand

## Slide 32: Inverse transfer function
The embedded processing unit of the sensor device applies the inverse of F(.), F-1(.), to convert the electrical signal generated by the sensing technique to output measurements in the same domain of the measurand.
Sensor
Approximated inverse transfer function
F-1(y)
Measurand
(e.g., glucose concentration)
Measurement (e.g., current in pA)
Glucose concentration mg/dl

## Slide 33: Common transfer function approximations
Linear transfer function approximation:
The offset, b, is the output value of the sensor when no measurand is applied. It is the component of the output signal which does not depend on the measurand.
The slope, a, represents the sensor sensitivity, i.e., the change in the sensor output resulting from a unit change in the input measurand.
Non-linear transfer function approximations:
Logarithmic function | 
Exponential function | 
Power function | 
Second-order polynomial function | 
Third-order polynomial function | 
Logarithmic function | 
Exponential function | 
Power function | 
Second-order polynomial function | 
Third-order polynomial function | 

## Slide 34: Calibration
The transfer function parameters, e.g., a and b in the linear transfer function, are in general unknown and must be estimated. This process is called calibration.
If a set of known (x,y) paired values are available, estimation of transfer function parameters can be performed by the least square approach.
For linear transfer function:
Two-point calibration
Multiple-point calibration

## Slide 35: Embedded processing unit
The embedded processing unit is responsible for:
  Performing local processing of the measurements
  Storing some data in a local memory
  Sending the data to another device (e.g., a display device or a gateway device for IoT systems)
Sensor
Embedded processing unit
Receiver device / Gateway device (for IoT devices)
Three main operations:
Denoising
Calibration
Fault detection

## Slide 36: Denoising
Raw signals are often noisy: the signal of interest is corrupted by noise
Example: high-frequency noise on a CGM signal

## Slide 37: Denoising
Raw signals are often noisy: the signal of interest is corrupted by noise
Example: low-frequency noise (e.g., due to movement) on an ECG signal
Clean ECG signal
ECG signal corrupted by low-freq. noise

## Slide 38: Denoising
The noise can be filtered by applying a denoising algorithm exploiting a digital filter.
Denoising algorithm
(digital filter)
Noisy signal
Clean signal
Non-causal filters: the filter output depends on past and future inputs  for offline use only
Causal filters: the filter output depends only on past and present inputs  suitable for online application

## Slide 39: Fault detection
Detection of temporary sensor errors also called faults or artifacts
Example: pressure-induced signal suppression in a CGM signal
Fault detection algorithm
Measured signal
Reconstructed signal
Signal reconstruction algorithm
(e.g. interpolation)
Artifact suppressed

## Slide 40: Communication systems
Wearable devices require a data transmission systems that wirelessly communicate between the sensing device and the receiver and/or gateway device (it can be the smartphone).
Bluetooth communication is the primary choice because of its simplicity, low operation cost, compatibility with new smartphone technology, and simple hardware. Drawback: it requires high power consumption.
Near field communication (NFC): a wireless communication technology that operates at a short distance (about 4 cm) and does not need the energy from the battery to transfer the data.
NFC tags can directly communicate with the smartphone equipped with the NFC modules.
Bypassing the need for batteries, NFC enables ultrathin, ultraminiaturized designs for sensors, like a temporary transfer tattoo.
BONUS

## Slide 41: The architecture of wearable IoT systems
1. Sensing layer
2. Network layer
3. Data storing and processing layer
4. Application layer
Alerts
Decision-support
Event log
Self-management
Data analyses
Modeling
Signal processing
Gateway device
WiFi,
Router
Internet
Sensors
Activity tracker
Blood pressure sensor
ECG sensor
Glucose sensor
GSM
Database

## Slide 42: 1. Sensing layer
This layer includes sensors, which are responsible for:
  Collecting the data either directly from the body through contact sensors or from peripheral sensors providing indirect information of body and behaviours
  Preparing the data for on-board analysis, closed-loop feedback or remote transmission for more comprehensive analysis and storing
ECG sensor
Glucose sensor
Contact sensors
Peripheral sensors
Step counter
Microphone and camera
Key components of IoT wearable sensors:
  Miniaturized sensor hardware
  Embedded processor with storage capabilities, power management and communication circuits
Typically, limited computing power and storage capabilities.

## Slide 43: 2. Network layer
Fore more comprehensive data analysis and storage, the sensor layer need to transmit the data to cloud computing servers (remotely-located computing machines).
This transmission can be mediated by a companion device that serves as a gateway, i.e., it enables the information to flow from sensors to the cloud server.
  Example of gateway devices: smartphones, tablets and PCs.
The communication must be secured through:
    Authentication protocols for writing, modifying and accessing the data
    Security and privacy protocols for preserving the data
Sensor
Gateway device
Cloud server
Bluetooth
WiFi
GSM

## Slide 44: 3. Data storing and processing layer
Layer responsible for data storing and complex processing
Key components:
  Database for structured data storing
    Data organized based on a specific data model (relational, hierarchical, object-oriented,…)
  Algorithms for data processing
    Statistical analysis
    Feature extraction (e.g., heart rate, calories expenditure, …)
    Signal processing (e.g., filtering, artifact removal, …)
    Data mining (e.g., cluster analysis, anomaly detection, sequential pattern extraction, …)
    Data modeling (knowledge-based, data-driven or hybrid approaches)
    Artificial intelligence (knowledge representation, machine learning, natural language processing, machine perception, automated planning, …)

## Slide 45: 4. Application layer
Layer responsible for visualizing raw or processed data to end-users and for providing decision support to them through a software application
Possible end-users: patients, physicians, caregivers, analysts, data managers, administrators, etc.
Possible types of applications:
Mobile app
Desktop app
Web app
App installed in a computer. It can work offline.
App that runs on the browser.
Installation not required. Internet access required.
App installed in a mobile device (e.g. smartphone, smartwatch, tablet, etc.)

## Slide 46: Appendix: Data source characteristics

## Slide 47: Focus on data – Calories
Calories: a list with 1 data point exactly every minute
Fields meaning:
  time: timestamp (hh:mm:ss format) of the calory entry
  value: the calories spent during the last 1 minute (rest + active)
[
  {
  "time": "00:00:00",
  "value": "1.29"
  },
  {
  "time": "00:01:00",
  "value": "1.29"
  },
  {
  "time": "00:02:00",
  "value": "1.29"
  },
  {
  "time": "00:03:00",
  "value": "1.55"
  },
  {
  "time": "00:04:00",
  "value": "1.29"
  },
  {
  "time": "00:05:00",
  "value": "1.29"
  },
  {
  "time": "00:06:00",
  "value": "1.29"
  },
  ...
]

## Slide 48: Focus on data – Distance
Distance: a list with 1 data point every minute or more depending on user’s movement
Fields meaning:
  time: timestamp (hh:mm:ss format) of the distance entry
  value: the distance travelled between the last timestamp and now (in cm)
[
{
"time": "00:03:00",
"value": "0"
},
{
"time": "00:50:00",
"value": "0"
},
{
"time": "01:28:00",
"value": "0"
},
{
"time": "01:32:00",
"value": "1440"
},
{
"time": "01:33:00",
"value": "3110"
},
{
"time": "01:34:00",
"value": "0"
},
{
"time": "01:53:00",
"value": "0"
},
{
"time": "03:09:00",
"value": "0"
},
...
]

## Slide 49: Focus on data – Steps
Steps: a list with 1 data point every minute or more depending on user’s movement
Fields meaning:
  time: timestamp (hh:mm:ss format) of the step entry
  value: the number of steps done between the last timestamp and now
Note: there can be days without data
[
{
"time": "00:07:00",
"value": "0"
},
{
"time": "00:10:00",
"value": "0"
},
{
"time": "01:57:00",
"value": "0"
},
{
"time": "02:29:00",
"value": "0"
},
{
"time": "02:40:00",
"value": "0"
},
{
"time": "02:57:00",
"value": "11"
},
{
"time": "02:58:00",
"value": "18"
},
{
"time": "02:59:00",
"value": "0"
}, ...
]

## Slide 50: Heart rate
Heart rate: a list with 1 data point every 5 seconds.
Fields meaning:
  time: timestamp (hh:mm:ss format) of the heart rate entry
  value: the heart rate at the current time (in bpm)
  confidence: a number from 0 to 3 defining how much the reading is reliable (this problably depends on user movement, etc. You can ignore it)
Note: there can be missing values
[
{
"time": "03:10:19",
"value": 58,
"confidence": 1
},
{
"time": "03:10:24",
"value": 54,
"confidence": 3
},
{
"time": "03:10:29",
"value": 52,
"confidence": 3
},
{
"time": "03:10:44",
"value": 52,
"confidence": 3
},
{
"time": "03:10:54",
"value": 53,
"confidence": 3
},
{
"time": "03:11:09",
"value": 54,
"confidence": 3
},
{
"time": "03:11:19",
"value": 53,
"confidence": 3
}, ...
]

## Slide 51: Resting heart rate
Resting heart rate: an entry per day.
Fields meaning:
  time: timestamp (hh:mm:ss format) of the resting heart rate entry (always 00:00:00)
  value: the estimated resting heart rate at the current time (in bpm)
  error: the estimate error (in %) (Can be ignored)
Note: there can be days without data
{
"time": "00:00:00",
"value": 51.35,
"error": 17.86
}

## Slide 52: Exercise
Exercise: a list with 1 entry for each recorded exercise session
Fields meaning:
  activityName: the type of entry (“Corsa”: Run, “Bici”: bike, “Camminata”: Walk)
  averageHeartRate: the average heart rate during the session
  calories: the active calories spent during the session
  distance: the distance covered during the session
  distanceUnit: the unit of the “distance” field
  duration: the total duration of the session (in ms)
  activeDuration: the total active duration of the session (in ms)
  steps: the total number of steps done during the session (if “Corsa” or ”Camminata”)
  logType: how this record was logged (“tracker” or “auto_detected”)
  heartRateZones: a list that summarizes the number of minutes and calories spent in specific heart rate ranges (from min to max)
  speed: the average speed (in km/h)
  vo2max: the estimated running VO2Max (only for “Corsa” activities)
  elevationGain: the elevation gained during the session
  time: the starting time of the session (format hh:mm:ss)
Note: different activity types can have (slightly) different fields
Note: there can be days without data
[
{
"activityName": "Corsa",
"averageHeartRate": 143,
"calories": 727,
"distance": 8.474117,
"distanceUnit": "Kilometer",
"duration": 3.548E+6,
"activeDuration": 3.214E+6,
"steps": 8604,
"logType": "tracker",
"heartRateZones": [
{
"name": "Fuori zona",
"min": 30,
"max": 110,
"minutes": 3,
"caloriesOut": 30.146219999999971
},
{
"name": "Grassi bruciati",
"min": 110,
"max": 137,
"minutes": 8,
"caloriesOut": 97.13782
},
{
"name": "Attivita aerobica",
"min": 137,
"max": 170,
"minutes": 48,
"caloriesOut": 659.73843
},
{
"name": "Picco",
"min": 170,
"max": 220,
"minutes": 0,
"caloriesOut": 0
}
],
"speed": 9.4981754822650917,
"vo2Max": {
"vo2Max": 48.77853
},
"elevationGain": 28.042,
"time": "11:22:43"
},
{
"activityName": "Bici",
"averageHeartRate": 150,
"calories": 3893,
"distance": 140.143695,
"distanceUnit": "Kilometer",
"duration": 2.0901E+7,
"activeDuration": 1.756E+7,
"logType": "tracker",
"heartRateZones": [
{
"name": "Fuori zona",
"min": 30,
"max": 110,
"minutes": 28,
"caloriesOut": 214.83527999999842
},
{
"name": "Grassi bruciati",
"min": 110,
"max": 136,
"minutes": 64,
"caloriesOut": 664.55027999999993
},
{
"name": "Attivita aerobica",
"min": 136,
"max": 170,
"minutes": 232,
"caloriesOut": 3167.1500100000012
},
{
"name": "Picco",
"min": 170,
"max": 220,
"minutes": 24,
"caloriesOut": 377.76059999999995
}
],
"speed": 28.734978587699313,
"elevationGain": 663.55,
"time": "11:03:54"
},...
]

## Slide 53: Sleep
Sleep: a list with 1 entry for each sleep session
Fields meaning:
  dateOfSleep: day associated to the sleep entry (MM-DD format)
  startTime: the starting timestamp of the sleep (MM-DD hh:mm:ss format)
  endTime: the ending timestamp of the sleep (MM-DD hh:mm:ss format)
  duration: the duration of the sleep session (in ms)
  minutesToFallAsleep: the number of minutes spent to fall asleep
  minutesAsleep: the number of minutes asleep during the sleep entry
  minutesAwake: the number of minutes awake during the sleep entry
  minutesAfterWakeup: the number of minutes the user spent in bed after waking up
  efficiency: the estimated sleep efficiency (from 0 to 100)
  logType: how the entry was logged (always “auto_detected”)
  mainSleep: a boolean indicating if the entry corresponds to the main sleep session
  levels: a summary of the sleeping stages (deep/wake/light/rem/restless) during the sleep session
  data: the stage profile of the sleep session. Each entry contains the starting timestamp, the level, and how much it lasted.
Note: there can be days without data
[
{
"dateOfSleep": "02-14",
"startTime": "02-13 22:44:00",
"endTime": "02-14 06:36:30",
"duration": 2.832E+7,
"minutesToFallAsleep": 0,
"minutesAsleep": 429,
"minutesAwake": 43,
"minutesAfterWakeup": 3,
"timeInBed": 472,
"efficiency": 96,
"logType": "auto_detected", ”mainSleep": true,
"levels": {
"summary": {
"deep": {
"count": 5,
"minutes": 106,
"thirtyDayAvgMinutes": 86
},
"wake": {
"count": 34,
"minutes": 43,
"thirtyDayAvgMinutes": 46
},
"light": {
"count": 28,
"minutes": 235,
"thirtyDayAvgMinutes": 281
},
"rem": {
"count": 11,
"minutes": 88,
"thirtyDayAvgMinutes": 101
}
},
"data": [
{
"dateTime": "02-13 22:44:00",
"level": "wake",
"seconds": 30
},
{
"dateTime": "02-13 22:44:30",
"level": "light",
"seconds": 150
},
{
"dateTime": "02-13 22:47:00",
"level": "deep",
"seconds": 1860
}, ... },
...
]